## Slide 1
Library QC for ONT metagenomics
Rapid Response – Updated January 2026

## Slide 2: Library Quality Control (QC)
The purpose of QC is to ensure that there is sufficient and good quality library for sequencing.
Quantity (concentration)
  Amount of library is measured using Qubit (or other fluorescence based method)
Quality (length)
  Measured using capillary electrophoresis (Tapestation, Bioanalyzer, Fragment Analyzer)
  If capillary electrophoresis is unavailable, regular gel electrophoresis may be used with a diagnostic PCR

## Slide 3: Overview of Library QC
Final QC
passed?
Sample A
Sample B
Water control A
Pool
Individual library QC passed?
Individual library QC passed?
Individual library QC passed?
YES
YES
YES
YES
QC 1
QC 2
Notes: Two moments during lib prep to do QC: individual library and pooled library

## Slide 4: Library prep QC
Concentration
Length
Remove adapter dimers
vs
vs
or
vs

## Slide 5: Assessing concentration is important for even pooling and accurate loading
Why is an accurate concentration important?
Individual samples: to pool equal amounts of different samples together
  Aim for an even amount of sequencing data per sample
Final library: The total amount of final library needs to be within a specific range*
  For the MinION 10.4 flow cell, this range is 10-50 fmol
    We have found that 100-150 fmol works best
*Nanopore is more forgiving than Illumina in accommodating a wider range of input loading amounts and we find that loading more than 50 fmol works well

## Slide 6: Library concentration is best assessed using fluorometry like Qubit
Fluorometry
(e.g. Qubit)
Detects fluorescent dyes that only bind DNA
✓
Use for concentration
Notes: RNA, guanidine

## Slide 7: Library concentration is best assessed using fluorometry like Qubit
Fluorometry
(e.g. Qubit)
Detects fluorescent dyes that only bind DNA
✓
Use for concentration
Electrophoresis
(e.g. QSep, Tapestation, Bioanalyzer, gel)
Compare intensity of DNA bands to standards
Especially useful to determine concentration in certain fragment length range
✓
Use for size profiling (“quality”) and molar concentrations (for pooling)
Notes: RNA, guanidine

## Slide 8: Bioanalyzer or Tapestation may be used for pooling, if necessary
Fluorometry
(e.g. Qubit)
Detects fluorescent dyes that only bind DNA
✓
Use for concentration
Compare intensity of DNA bands to standards
Especially useful to determine concentration in certain fragment length range
✓
Use for size profiling (“quality”) and molar concentrations (for pooling)
Electrophoresis
(e.g. QSep, Tapestation, Bioanalyzer, gel)
Notes: RNA, guanidine

## Slide 9: Nanodrop is not recommended for assessing concentration
Spectrophotometry
(e.g. Nanodrop)
DNA absorbs light most strongly at 260 nm
Level of absorbance ~ concentration of DNA
BUT other molecules (RNA, proteins) & contaminants also absorb light at 260 nm, causing up to 10-fold overestimation of the DNA concentration!
✓
Use for detecting contaminants
✘
Do not use for concentration
Notes: RNA, guanidine

## Slide 10: How to use the Qubit
Qubit dsDNA High Sensitivity (HS) Assay kit
dsDNA HS Reagent (200x concentrate)
dsDNA HS Buffer
dsDNA HS Standard #1
dsDNA HS Standard #2
Prepare Working Solution:
  Diluting HS reagent 1:200 in HS buffer
Dilute in Working Solution:
  Standards: 10 μl standard + 190 μl Working Solution
  Samples: 1 μl sample + 199 μl Working Solution
Vortex 2-3 seconds before running standards/samples

## Slide 11: How to use the Qubit
Select dsDNA HS workflow
Run the 2 standards
Enter sample volume added to tube
Run sample
Top value = original conc.
Bottom value = diluted conc.
Qubit dsDNA High Sensitivity (HS) Assay kit

## Slide 12: Library prep QC
Concentration
Length
Remove adapter dimers
vs
vs
or
vs
Notes: Update to protocol: Ct <15: dilute 1/10, rest use undiluted

## Slide 13: Assessing length can provide insight to the quality of the library: Is the library good enough to load?
Why is it important to assess library size distribution?
Depending on sample quality, there can be huge variations in the library size (100s of bases to 1000s of bases)
Non-specific products can sometimes be formed
  Primer/oligo dimers when high input cDNA concentrations are used
  Long concatemers that can block pores and reduce sequencing output (~10kb or more)
Fragment length distribution impacts molarity calculations for pooling and sequencer loading
  Convert from ng/μl (Qubit) + Length bp to nM (fmol/µL)

## Slide 14: Library size is assessed with electrophoresis methods
Gel electrophoresis
Capillary electrophoresis
E.g. Tapestation, Bioanalyzer
PCR amplify an aliquot of the library using “universal” primers
Gel electrophoresis of ladder and samples
Add buffer to aliquot of library
Run machine
Takes ~ 4 hrs
Does not require specialized equipment
Tends to slightly overestimate library size
Tapestation ~ 10-20 minutes
Bioanalyzer ~ 45 minutes
No PCR bias
Can calculate concentration in nM
How to assess library size distribution?

## Slide 15: Diagnostic PCR method
Take 1 uL of the individual libraries or final pool and perform a standard PCR reaction for 20 cycles
Note: the PCR product will only be used for diagnostic QC, not for sequencing
For each sample or pool, setup the amplification reaction as follows:
Component | Volume (µL)
LongAmp Master Mix (2x) | 25
Nuclease-free Water | 23.5
RLB-PCR Primer* (10µM) | 0.5
DNA Library Template | 1
Total Volume | 50
Use the following cycling conditions:
Thermocycler (lid at 105˚C) | Cycles
95 °C for 45s | 1
95 ˚C for 15s | 20
56 ˚C for 15s | 
65 °C for 5 min | 
65˚C for 10 mins | 1
Hold at 10˚C | ∞
Run PCR amplified product on a gel to visualize and estimate the library size.
CZ Biohub Protocols: Nanopore mNGS
*This primer is only for diagnostic PCR

## Slide 16: Diagnostic PCR results compared to Tapestation
1 2 3 Neg
1
2
3
~1700bp
~900bp
~1200bp

## Slide 17: Library size is assessed with electrophoresis methods
Gel electrophoresis
Capillary electrophoresis
E.g. Tapestation, Bioanalyzer
PCR amplify an aliquot of the library using “universal” primers
Gel electrophoresis of ladder and samples
Add buffer to aliquot of library
Run machine
Takes ~ 4 hrs
Does not require specialized equipment
Tapestation ~ 10-20 minutes
Bioanalyzer ~ 45 minutes
No PCR bias
Can calculate concentration in nM
How to assess library size distribution?

## Slide 18: Tapestation genomic DNA kit can be used to assess library size
Tapestation
Genomic DNA kit is most suitable since fragments can be 1000s of bases long
✓

## Slide 19: Look out for ~10kb concatemer artifacts - they will clog the flow cell pores
Tapestation
Genomic DNA kit is most suitable since fragments can be 1000s of bases long
✘
Concatemer peak at ~10kb, with no other peaks*
*Consider the input going in, and the library size you would expect. In this case, a 10kb peak was seen for RNA input, which is much longer than the fragment size we’d expect. It turns out that the 10kb peak was a concatemer artifact created during barcoding PCR.

## Slide 20: Library size
Checking library size is important for troubleshooting
It is recommended to routinely check library size when establishing the protocol or expanding to new sample types
If results are consistent and sample type / source are constant, it is possible for the library size to be estimated
Size may be evaluated by:
  Diagnostic PCR + gel electrophoresis
  QSep / Tapestation / Bioanalyzer / Fragment analyzer
Size will be used for concentration calculation

## Slide 21: Library prep QC
Concentration
Length
Remove adapter dimers
vs
vs
or
vs
Notes: Update to protocol: Ct <15: dilute 1/10, rest use undiluted

## Slide 22: Library prep QC - Clean up primer/oligo dimers
Primer/oligo dimers:
  Occur when primers and unused 9N and TSO oligos interact during the PCR amplification
  Will not be informative if sequenced
Primer/oligo dimer peak

## Slide 23: SPRI beads may be used to remove adapters and reagents
SPRIselect or AMPure XP SPRI select beads (paramagnetic)
The longer a fragment, the better it will bind to the beads
The fewer beads you add, the more short fragments you’ll lose
Perform size selection using these beads to remove adapter dimers
Volume beads
Volume reaction
SPRI ratio =
Example: for a 1.8x SPRI clean-up:
Add 18 μl of SPRI beads to 10 μl of sample
1. Bind DNA to magnetic beads; 2. Separate beads from contaminants;
3. Wash magnetic beads with 85% ethanol to remove contaminants;
4. Elute DNA from magnetic beads; 5. Transfer sample to new tube

## Slide 24: Library prep QC - Clean up adapter dimers
AMPure XP SPRI select beads (paramagnetic)
Size selection using SPRI beads to remove adapter dimers
Samples can first be pooled and then a 0.6x SPRI clean-up can be performed

## Slide 25: Library prep QC
Final QC
passed?
Sample A
Sample B
Water control A
Pool
Individual library QC passed?
Individual library QC passed?
Individual library QC passed?
YES
YES
YES
How to fix samples with QC issues?
Concentration too low -> amplify or restart
Adapter dimers -> perform additional cleanups
Concatemer artifact observed -> restart
NO
NO
YES
Fix it
Fix it
Notes: Two moments during lib prep to do QC: individual library and pooled library

## Slide 26: How much concentration do I really need?
Final pool
  Need 100-150 fmol to load the flow cell
  = 11ul of 10-15nM
Individual samples:
  Ideally individual samples would also have 10-15nM
    (7-10ng/ul for 1000bp library)
    (13-20ng/ul for 2000bp library)
  Total mass (ng) per sample added to pool will depend on:
    library size
    # of samples in the pool

## Slide 27: Using the Tapestation

## Slide 28: Choosing the right assay according to your starting material
DNA
200 to >60000 bp, 10 – 100 ng/µL
RNA
500 – 10000 pg/µL
Sequencing libraries
100 - 5,000 bp, 10 - 1,000 pg/µL
Note, sensitivity is guided by input concentration.

## Slide 29: How to analyse a TapeStation run?
Ensure that all marker peaks (‘Lower’ and ‘Upper’) are properly assigned. Concentration values are calculated using the area of the sample peak compared to the known concentration of the upper marker.

## Slide 30: How to analyse a TapeStation run?
Ensure that all marker peaks (‘Lower’ and ‘Upper’) are properly assigned. Concentration values are calculated using the area of the sample peak compared to the known concentration of the upper marker.
Ensure that sample peaks are properly integrated by manually adjusting the peak when necessary.

## Slide 31
This is the “Electropherogram” view setting. If you see this after the TS analysis ends, click on “Region”

## Slide 32
Click on “Region Settings”

## Slide 33
Click on this line to add the peak range of interest

## Slide 34
Enter the “from” and “to” range here. We commonly use the 200-1000 bp range
By selecting “assay file” this range will apply to all future assays. But you can also specify the range for a sample or file.
Notes: You can set multiple peak ranges as well

## Slide 35
Click “Apply”

## Slide 36
The “Region Molarity” concentration is now based on the 200-1000 bp range selected. This is the TS value that we use for our sample quantification

## Slide 37
You can click and drag the peak range vertical line to make adjustments. In this case, we shifted the line from 200 bp to 132 bp to take into account shorter fragments. Region molarity is automatically recalculated.

## Slide 38
When you click on file this window will open. Click on “Create report” on the left panel in the purple box. That will create the report in pdf which is now you can save by clicking on “Save As…” in brown box. This will create the pdf report wich is amenable to view in other computers

## Slide 39
Save and now your report is ready to share as a pdf!